## Slide 1: The HV Protection Boards For The RICH Detectors Of LHCb
Claudio Arnaboldi, Tito Bellunato, Paola Gobbo, Davide Luigi Perego and Gianluigi Pessina
Istituto Nazionale di Fisica Nucleare (INFN)
Università degli Studi di Milano Bicocca - Dipartimento di Fisica
P.za della Scienza 3, 20126 Milano, Italy
On behalf of the LHCb RICH Collaboration

## Slide 2: ABSTRACT
We present the circuit protection system for the monitoring of the High Voltage bias for the Hybrid Photon Detectors (HPDs), of the Ring Imaging CHerenkov (RICH) detectors of LHCb. The protection system buffers the voltage lines, attenuated by 12800 V/V, in normal operating conditions and limits the output voltage excursion to the ADC monitoring system to a safe range (between +5 V and -2.5 V) in case of discharge. The circuit is designed to be radiation tolerant. The protection system has been fully characterized with radiation and temperature under the whole expected working conditions of LHCb. Results have shown that the developed protection boards are fully adequate for the whole LHCb lifetime.

## Slide 3: The detector
RICH1 and RICH2 are Ring Imaging CHerenkov detectors at LHCb. Photons generated inside the detector are converted to electrons in 488 Hybrid Photon Detectors (HPD). The HPDs are distributed in columns to form a close-packed configuration:
1/4 of
RICH2
HPD
-19.7 KV
-16.4 KV
-20 KV
GND
Every HPD needs 3 High Voltages, HV, that are responsible of the acceleration, focalization and demagnification of the photo-electrons toward the pixel chip, at about 100 V above GND potential.
100 V
Column
The managing of the HVs is made with PCBs embedded in Silicone rubber.

## Slide 4: HV distribution scheme for one column
The -20 KV feeds the Splitter board where the other 2 voltages (-19.7 KV and -16.4 KV) are generated. The Splitter boards also provide biasing to 2 HPDs.
The 3 HV's then feed the Intermediate boards that bias a further pair of HPDs, and feed the HVs to the next boards in the chain.
The last board of the half-column is the Monitoring board that, as well as biasing a pair of HPDs, generates 3 copies of the HVs attenuated by about 12700 V/V.
The attenuated HVs can be readout by the monitoring system.
-19.7kV
-20 kV
-16.4kV
A
B
C
A
B
C
A
B
C
A
B
C
5G
One half-column of the RICH
Splitter
Monitoring
Intermediate
Every column is divided in to 2 half-columns for its HV distribution. A set of boards manages the HV. Every board serves a pair of HPDs.

## Slide 5: The electrical details of the Monitoring Board
Monitoring lines
To ELMB
(Embedded Local Monitor Board )
Voltage divider (0V , -1.6V)
Surge arrester -90 V
The 3 attenuated voltages to be monitored must be protected against any possible discharge. A Surge Arrestor is included on the Monitoring Board to limit to a safe 90 V. This is not enough to prevent permanent damage to the ELMB, the multi-channel ADC system used at LHCb.
ELMB safety range
-2V - +5V
We need further protection to be put between the Monitoring board and the ELMB.
We need to add
protection

## Slide 6: Introducing the Protection Board
-20 KV
-19.7 KV
-16.4 KV
5 G
392 K
X 2
The Protection Board is a buffering system able to copy the voltages to be monitored while assuring the safe condition under any discharge.
The specifications the Protection Boards are able to satisfy are:
Unity gain inverting buffering (the ELMB has a larger operating positive range);
Safe margin between -2.5 V and 5 V against 150 V input discharge;
7 channels per board to manage 6 signals from 2 half-columns and the bias voltage of the Silicon pixel arrays of the same column;
Radiation tolerance against the working location (within the radiation area of LHCb), 9.61011 n/cm2, 700 rad;

## Slide 7: Protection Board schematic diagram
150 V
We selected CMOS OAs (Operational Amplifier) having protecting diodes towards the rails at the inputs. 66 K and 33 K resistors have been put in series with the inputs to limit the input current through the diodes to safe limit in case of discharge. 2 unity gain buffer are used to avoid any dropout effect across the series resistors, since their negligible input current.
Supply voltages are +4.7 V, -4.7 V and -2.1 V. The OAs are biased in such a way the 2 differential output voltages are limited between -2.1 V and 4.7 V under any discharge conditions.
The circuit is differential input - differential output with a 15 Hz low pass filter at the output.
Accuracy of resistors used was 0.1 %.

## Slide 8: Protection Board Characterization
-150 V
To the Protecting board inputs
Discharge
A very accurate characterization procedure has been developed. A remote controlled relays system allow to inject signals of known amplitude to simulate discharges.
The testing/characterization system was able to manage 5 boards at a time.
Characterization was done with temperature (inside an environment chamber) and under irradiation (in a reactor).

## Slide 9: Set-up of the characterization system
PC with MATLAB
150 V Power supply
Multi-channel multimeter
Environment chamber
or
Radiation environment
Scope
5 boards
Digital driver
Switching relays 1
Switching relays 2
GPIB
RS-232
GPIB
Tested parameters against temperature and irradiation:
DC Gain;
Linearity
Offset;
Drift;
Discharge;
Voltage Supplies;

## Slide 10: Results example
Repeated input pattern
Linear regime
Discharge simulation
Linear regime
Discharge
Every characterization test lasted about one day.
The input pattern was repeated periodically, every about 1.5 hours.
The selected voltage values allowed to fully characterize the channels under test.
The outputs signals were measured in response to the input pattern to study the parameters under investigation.
The same set-up was used vs Temperature and Irradiation.
Output from one channel

## Slide 11: Temperature dependance
OA output offset Gain
(0.040±0.730) mV -(0.999±0.002)
(0.264±4.640) μV/°C -(0.368±6.280) ppm/°C
ΔT(LHCb) ≈ ±5 °C
Maximum effect: 0.003 %
35 boards with 245 individual channels has been fully tested. No failures have been observed with respect to the simulated discharges.
The temperature variation of the gain and the offset has shown that the chosen design and technology is fully compliant with the expected specifications at LHCb.
Average result

## Slide 12: Radiation tolerance measurements
Irradiation test has been made at the reactor TRIGA MARK II situated in Pavia, IT
The Protection boards will be located in the radiation areaof LHCb where the maximum expected level of radiation after 10 years of data taking is:
γ dose : 700 Rad
Neutron fluence: 9.6 E+11 1MeV eq ncm-2
4·1013 ncm-2 40 x LHCb
~43 kRad 60 x LHCb
We irradiate 3 boards, 6 h each. After this period the level of irradiation applied was:
We gave to the Protection Boards a much larger irradiation dose with respect to that expected during the LHCb lifetime.

## Slide 13: Irradiation results
Off(1013) – Off(0) = (0.318 ± 0.590) mV
G(1013) – G(0) = -(0.071 ± 0.261) %
The maximum spread was:
it was not possible to find a correlation between the irradiation level and the parameters variations;
The effect of the doses gave a negligible effect - especially if we limit the irradiation to that expected in LHCb, the red line in the 2 graphs.
The effect of the doses gave a negligible effect. If we limit the irradiation to that expected at LHCb, the red line in the 2 graphs, the effect is even more marginal.
Remarks:

## Slide 14: Conclusions
The monitoring of the HV bias of the HPDs of the RICH1 an RICH2 at LHCb has been performed with a full level of safety;
The protecting circuit that has been designed, built and already in use, has shown outstanding level of accuracy in both the static and dynamic parameters;
Radiation tolerance was found to be more than adequate for the whole lifetime of LHCb.